## Slide 1
NovaReader-Flower System
Design Introduction
3 Jan 23
MDC-Simon
Ver 0.1

## Slide 2
NovaReader- Flower Function Block
Architecture Diagram
Main Control Board
Auto Machine Dimension (Draft)
Auto Machine Specification (Draft)
Cartridge Drawing (Draft)
Sensor PCB

## Slide 3: NovaReader- Flower Function Block
Display
LCD + Touch Panel
Barcode Scanner
Biosensor (NS01B)
Memory
Interface
MCU(Micro Control Unit) ARM Cortex M3 32bit
Flower System
Cartridge install stage
Cartridge install/remove sensor
Lid open/close sensor
Probe contact up/down stage
Cartridge rotate control
Tube push/pull control part
Housing mounting bracket and base
Power supply (External adapter)
Motion System
Communication
DC 12V 1A

## Slide 4: Architecture Diagram (Main Control)
Display
LCD + Touch Panel
Barcode Scanner
FEA(Front End Amplifier)
Biosensor (NS01B)
Memory
TF Card
NOR
Flash
Interface
MCU(Micro Control Unit) ARM Cortex M3 32bit
Ethernet
USB
MCU
Motion System

## Slide 5: Main control board

## Slide 6: Auto-Machine Dimension (Draft)
Dimension: 300mm W x 200mm D x 250mm H
Weight: 3~4KG

## Slide 7: Auto-Machine Specifications
Power Supply (Main Power) | 100~240VAC 50/60HZ for motion parts | 
Power Supply (Sub Power) | DC12V 1A for Flower system operation | 
Operating environment | Indoor use 10℃~40℃ 10% ~ 60% RH(with condensing) 0m(1.01bar) ~ 2000m(0.763bar) IP 20 | 
Storage environment | 10℃~40℃ 10% ~ 60% RH(with condensing) 0m(1.01bar) ~ 2000m(0.763bar) | 
For safety, the power supply must be an isolated power system

## Slide 8: Cartridge Drawing

## Slide 9: Sensor PCB
8 Chips
1 Chip